## Slide 1
Accurate profile provided for Φ = 2.4 mm, P = 4 bar
Profiles linearly extrapolated for P = {3, 2, 1} bar
From the datasheet:
AP Inlet Valve Diagrams

## Slide 2
Simscape Model

## Slide 3
Simscape Model
Inlet Valve model
Time constant = 10 ms
Preload spring offset:
Inlet Flow Controller Integrator + Adaptive FeedForward

## Slide 4
AP Inlet Valve Responses

## Slide 5
AP Inlet Valve Responses

## Slide 6
AP Inlet Valve Responses

## Slide 7
AP Inlet Valve Responses

## Slide 8
AP Inlet Valve Responses
🛑 close to 100%

## Slide 9
AP Inlet Valve Responses

## Slide 10
Tidal Volume Drop w/ Helmet

## Slide 11
FMEA
@ :
Inlet fully open
Outlet closed

## Slide 12
AP + CFB: Volume Control

## Slide 13
AP + CFB: Pressure Control

## Slide 14
CFB + CFB: Volume Control

## Slide 15
CFB + CFB: Pressure Control

## Slide 16
CP + CFB: Volume Control

## Slide 17
CP + CFB: Pressure Control

## Slide 18
Simscape Model w/ sensors

## Slide 19
Controller

## Slide 20
AP + CFB: Volume Control w/ sensors
@ 1 kHz

## Slide 21
AP + CFB: Pressure Control w/ feedback-filtering, input-shaping
Feedback Filtering
IIR 2nd order
5 Hz cutoff
25 dB attenuation
0.1 dB passband max ripple
Input Shaping
Minimum Jerk 3rd order
@ 1 kHz

## Slide 22
AP + CFB: Pressure Control w/ feedback-filtering input-shaping, gain-scheduling
Gain Scheduling
Different I+ and I- gains
@ 1 kHz

## Slide 23
AP + CFB: Pressure Control comparison w/ and w/o gain-scheduling

## Slide 24
AP + CFB: Pressure Control w/ feedback-filtering input-shaping, gain-scheduling
Feedback Filtering
IIR 2nd order
10 Hz cutoff
25 dB attenuation
0.1 dB passband max ripple
Input Shaping
Minimum Jerk 3rd order
@ 100 Hz

## Slide 25
AP + CFB: Pressure Control w/ latest tabled values
Feedback Filtering
IIR 3rd order
8 Hz cutoff
30 dB attenuation
0.5 dB passband max ripple
Input Shaping
Minimum Jerk 3rd order
Gain Scheduling is not required
@ 100 Hz

## Slide 26
CP + CFB: Pressure Control w/ latest tabled values
Feedback Filtering
IIR 3rd order
8 Hz cutoff
30 dB attenuation
0.5 dB passband max ripple
Input Shaping
Minimum Jerk 3rd order
Gain Scheduling is not required
@ 100 Hz

## Slide 27
Simscape Model (update)

## Slide 28
AP + CFB: Pressure Control w/ new pressure profiles from MVM

## Slide 29
AP + CFB: Pressure Control w/ new pressure profiles from MVM

## Slide 30
CP + CFB: Pressure Control w/ new pressure profiles from MVM